December 23, 2009

Kent, matt

Disconnected 7500 and uploaded new program to towertop system.

Swapped cards on all systems.

Downloaded all hobos except uppermost one.

Installed new hobo batteries in all hobos except uppermost one.

Found that one of the hobos on hort. System was not working.

December 14, 2009

Matt, kent

Get the new charger system working.

December 12, 2009

Kent, matt

Swapped cards on subcan

Installed fresh HMP on subcan system.

Installed new battery charging system, but did not get it working yet

November 30, 2009

Kent, matt

Swapped all cards

November 10, 2009

Kent, Jon

Swapped all cards

Calibrated subcan 7500.

Zero h20 reading 0.56 mmol/mol (-29C), set to zero

Span h20 reading 6.05, set to 6.63C

Zero co2 reading 12ppm, set to zero

Span co2 reading 398, set to 403.85

November 7, 2009

Jon Martin, Kent

Calibrated 7500 and 7000 on towertop

Fixed Hobos on horizontal array that had not been working.

Found that subcan was not working.  Found that HMP was short-circuiting the logger causing a self calibration error.  Took HMP off the logger.

Downloaded hobs on vertical array.  Found that top one needs to have battery changed.

Forgot  to download towertop logger.

September 30, 2009

Matt, Kent

Swapped out bad Sub can CR5000

Installed a new CR5000 for horizontal array.

Swapped cards on profile and vertical loggers, but not for tower top logger

Downloaded all horizontal hobos.

August 7, 2009

Kent, Matt

Swapped out Horizontal CR5000 for a newer one.  Need to send the old one back for calibration.

Messed with Handar 6 again.  Must be the handar, even though it works in the lab.

Bees nest in the battery shed.

July 28, 2009

Changed all cards

Recalibrated the 7000- but still not right with the water channel

Downloaded Hobos on both Handar arrays

Installed new N2 tank

Got the CR5000 logger for the vertical array working.

Replaced batteries on three hobos.

Tried to get Handar 6 back up, but no luck

July 10, 2009

Some one changed cards

June 30, 2009

Chris, Kent

Calibrated closed path.

Tried to install handar 6, but no luck.

Installed Picarro flux analyzer

June 26, 2009

Kent, Jon M

Swapped cards

June, 19, 2009

Kent, Matt

Installed new N2 tank for closed path

Installed master on/off switch for whole site

Fiddled with handars

Removed handar 4 and 6 on horizontal

Removed handar 13, 15 on vertical array

June 10, 2009

Jon Martin, and etc…

Swapped cards on all loggers.

May 22, 2009

Kent, Jon

Noticed that closed path was not working. Restored power to the instrument and it seemed fine.

Swapped all cards

Downloaded hobos on tower and changed all except the upper most one to 4 minutes

Noted handars on tower

Handar Id		Handar #		Height

2			?				0.3m

6			“a”				1.50m

13			“o”				2.80m

15			“3”				6.4m

12			“1”				9.45m

4			“6”				13.25m

9			“f”				17.75m

5			“9”				22.95m

7			“c”				29.70m

Note:

15 and 12 are at the same height, but on different towers.

2 is mounted upside-down.

Moved the 2 lower most inlets of the profile system to a new location closer to the subcan system

Jon measured new bearings for the Handar network.   Compass did not have a magnetic declination set.

S4-S8=288

S5-S8=234

S5-S4=209

S10-S11=248

S6-S7=167

S6-S12=33

May 13, 2009

Jon,

Powered up the system

May 8th, 2009

Cory

The generator was not starting and battery voltage was very low so Cory pulled the plug on the site.

April, 30,2009

Kent, Matt, Jon

Swapped cards on subcan and towertop systems.

Swapped N2 tanks for closed path.  The tank only had 800psi in it so we need to take out another one soon.

Replaced S01 handar on hort array

Replaced first handar on Vert array

Downloaded Hobos on Hort array and changed them over to 4 min sampling interval

Renamed a few of the hobos

Made some notes about handars

Hobo	station		handar num	address		output label

69	S1		11		“j”		H1

99	S4		14		“e”		H4

125	S8		16		5		 H8

114	S5		17		6		H5

117	S6		19		0, but says “a” on Handar		H6

156	S12		21		“e”		H12

119	S9		22		“f”		H9

116	S7		18		9		H7

115	S10		2		3		H10

134	S11		3		5		H11

Bearing from S4 to S8 is 270 W (true)

Bearing from S5 to S8 is 225 SW

Bearing from S12 to S6 is 234 SW

Bearing from S10 to S11	270

April 21, 2009

Kent, Matt

Downloaded all handar data and co2 data

Moved Handar from S01 to second position on vert array.  H11 (from second position on vert array) was not responding so I brought that one in.  will take it back out.  Noticed that handar from first position was not working.  Will replace with a different one.

April 1, 2009

Not sure what happened

But did swap cards

March 4th, 2009

Kent, Garrett, Matt

Swapped all cards

Installed new N2 for closed path

Swapped filter for closed path.

Every thing else seemed fine.

Feb 12th, 2009

Kent

Replaced Handar 4 and 13.  They should be recording good data now.

Feb 4th, 2009

Kent, Chris, Jon Martin

Swapped all cards.

Downloaded all Hobos.  Replaced batteries on three hobos, and replaced one hobo (125).

Installed new N2 tank for closed path.

Found handar 4 and handar 13 not working.

Adjusted orientation of metOne wind sentry on towertop: now pointing to magnetic north.

Checked orientation of sonic at towertop: 350 deg magnetic bearing

Jan 14, 2009

Kent, Jon Martin

Fixed tree on guy wire and swapped card on towertop system.  Downloaded horizontal handars.  Found that 3 where not working.

Jan 5th 2009

Kent, Jon Martin

Changed cards in all loggers except the towertop.  Found that a treetop was under guy wire.

Changed N2 for closed path system.