January 11, 2011

Matt, Kent

Downloaded data cards and hobos, except hobos #0&1 (on small tower)

Replaced subcanopy 7500 with -0527, however, DPG was not working so will have to calibrate it tomorrow.

Replace Handar 19 with a new one and it seems to be working.

Propane at 35%

January 5, 2011

Kent, Matt

Replaced N tank

Replaced modem at subcanopy – timestamp off by 6 days.  Need to replace internal battery.

Found that 7500 on subcanopy is reading negative water channel.  Need to replace the 7500 with one that has fresh desiccants.

Brought in handar from horz array for testing and addressing

January 4, 2011

Kent, Matt

Replaced & relaunched hobo#7

Replaced batteries and relaunched hobos #2&3

Downloaded data from other hobos

Swapped cards

Checked modem at subcanopy, appears to have died

Handar #3 on the tower is not working.  That would be handar # 19.

Need to replace N tank

December 6, 2010

Kent, Matt

Replace intake filter on closed-path

Swapped cards

Downloaded hobos; #2,3, and 7 (bottom two and top one on tower) indicate  communication failures.

Removed a handar from horizontal network for use at MP

Propane at 50%

Noticed that the bottom 4 handars where not reporting data, power cycled them and they seemed to come back online.

November 12, 2010

Kent

Replaced 840 with 820 and calibrated it.  Also changed plumbing so that only ambient gas goes through dryers.

November 10, 2010

Kent, Matt

Replaced 840 with 820 on profile system.

Topped off oil in generator

Propane at 65%

November 8th, 2010

Kent, Matt

Swapped all cards,

Replaced N2 with full tank

Replaced half-hour tank for profile

Noticed that irga was not responding on profile

Downloaded hobos and renamed them from 0-7; replaced battery on hobo #2 (lowermost on tower)

November 4th, 2010

Kent, Matt

Replaced 325.68 with 341 tank.

Replaced dryers.

Swapped cards on profile system only.

Replaced N2 tank with one that has 400 psi.

October 15, 2010

Jon, Matt, Kent

Swapped filters on closed path and profile inlets

Swapped cards

Downloaded hobos except for the two on the mini tower-noticed 3rd (labeled  MF2_139) one is not working, brought in to be replaced. Update 101019 (kd) this one seems to work in lab, but it had lost its address and interval.

Also MF3_73 is bad, needs replacement

We need to fix labels!

Started generator manually

October 8, 2010

Swapped cards

Downloaded hobos

September 29, 2010

Jon Martin replaced profile pump.

Sept. 27, 2010

Kent, Matt

Noticed that profile pump is down.

Replaced broken drive belt, spark plugs.

Fired up generator.

Turned site back on.

Sept. 21-24, 2010

Matt

Diagnostics on generator.

Topped off propane tank.

Acquired drive belt and spark plugs from Cummins in Coburg

Sept. 15, 2010

Matt

Generator will not start.  Has spark & compression, cannot detect any gas odor at inlet line.  Shut down site to preserve batteries, turned off all tanks.

Propane still at 60%

Called CoEnergy for propane line service tech.

Sept. 8, 2010

Jon, Matt

Swapped card at tower top

Replaced battery and relaunched topmost tower hobo

Downloaded all hobo data

Swapped N2 tank, old one still has 400psi & was left at site

Took LAI measurements at all plots

Took phenology measurements

Topped off generator oil

Propane at 60%

Looks like hobo 4 needs to be renamed.

September 1, 2010

Kent, matt, jon

Swapped all cards,

N2 tank had about 800 psi left, left a full tank in shed.

Downloaded hobos- seems 4 has lost its name and the topmost one has communication failure.

Propane is at 65%.

July 30, 2010

Matt

Changed oil, oil filter, and air filter on generator

Launched hobos on short tower

July 29, 2010

Kent, matt

Calibrated 7000,  fixed leak on plumbing

Swapped cards on vert handar array, subcan, and towertop loggers.

July 28, 2010

Kent, Matt, Jon

Foggy, slight wind, slight rain.

Attempted to calibrate 7000.  Found that it may have had a leak in the plumbing, could not get a good calibration. Then thunder started up.  Jumped ship.

Matt replaced batteries on the first two hobos, but did not launch them.

Calibrated 7500 on subcan system.

Subcan calibration zero h20 reading -25C, set to zero

Span reading 4.34, set to 5

Zero co2 reading -1.76, set to zero

Span co2 reading 397, set to 403.07

July 21, 2010

Christoph

Dismantled horizontal hobo array

July 21, 2010

Kent, Matt

Swapped cards on everything except Hortizontal handar array.

Downloaded Hobos and replaced Hobo 5 with a new one.  Top Hobo and first and second  Hobo need batteries.  Both Hobos on short tower (not subcanopy) also need batts.

Calibrated 840 profile system.

Took down 7500 from tower.

Replaced N tank for 7000

Need to replace driers for profile system

Also need to replace the filters

Propane 70%

July 13, 2010

Matt

Checked system- realtime data had not refreshed since 7/9/10 18:00

12.6v at battery terminals, 10.6v at power terminal blocks, all good

Propane at 70%

A calibrated hobo is in the shack, ready for installation

Replaced CO2 tank CC7424@402.45ppm w/ CC91364@389.49ppm

July 1, 2010

Matt, Jon

Brought new N tank to site

Reinstalled power relay PCB

Maintained soil respiration chambers

Removed empty N tanks

Propane 78%

June 26, 2010

Matt

Manually started generator

Topped off generator oil

Fried power relay PCB, removed for repair

June 25, 2010

Kent, Matt

Downloaded cards

Replaced failed hobos

June 21, 2010

Matt

Filled propane tank

June 1, 2010

Kent, Matt

Download all cards,

Downloaded all Hobos-noticed that uppermost and 2nd from bottom are not working

April 5, 2010

Kent, Jon

Changed N2 on closed path

Swapped all cards

Downloaded HOBOS on tower -noticed that uppermost one is not working

Propane is at % 40

March 3, 2010

Kent, Matt

Can’t remember right now

January 25, 2010

Kent, Matt

Swapped out new 7500 Head  (0557) with new desiccants on subcanopy.  Seemed the water channel was way out.

Found generator running at about 11:40.  could not figure out why.

January 22, 2010

Kent, Matt, Jon

Installed new program in generator.  It should only run 10 hours now.  Found that 20amp breaker on generator could be bad.